## About
Form Number:
Title:
Revision Date:
Purpose:
Distribution:
DT-1695 (T-233)
LABORATORY TEST EQUIPMENT INSPECTION WORKBOOK
10-02
To record and report inspection results of equipment in bituminous materials, hydraulic cement, portland cement, and soils laboratories.
Original to Headquarters Materials and Tests; copies to Regional Materials and Tests, Project Supervisor.
TDOT E-Forms
Materials and Tests Electronic Forms
Microsoft Excel 97
Copyright 1984-1997 Microsoft Corporation
Tennessee Department of Transportation
Division of Materials and Tests
Field Operations
6601 Centennial Blvd.
Nashville, Tennessee 37243-0360

## DT-1695A
STATE OF TENNESSEE
DEPARTMENT OF TRANSPORTATION
DIVISION OF MATERIALS AND TESTS
6601 CENTENNIAL BLVD.
NASHVILLE, TENNESSEE 37243-0360
AGGREGATE TEST EQUIPMENT INSPECTION
Date
Location |  |  | County | 0 |  |  | Region | 0
Inspector
Equipment | Test Method | Requirement |  | Interval (Months) |  | Reference | Inspected | N/A
Bulk Density (Unit Weight) Measures | T19, C29 | Calibrate |  | 12 |  | AASHTO T19, ASTM C29
Mechanical Shakers |  | Ck. Sieving Thoroughness |  | 12 |  | AASHTO T27
Gen. Purpose Balances, Scales & Weights |  | Verify |  | 12 |  | AASHTO M231
Sieves, Coarse |  | Ck. Openings & Physical Condition |  | 6 |  | AASHTO M92, E11
Sieves, Fine |  | Ck. Physical Condition |  | 6 |  | AASHTO M92, E11
Ovens |  | Verify Temp. Settings |  | 4 |  | DMT-C/V-#3
Sulfate Oven | T104, C88 | Ck. Rate of Evaporation |  | 12 |  | AASHTO T104, DMT-C/V-#3
L.A. Machine | T96, C131 | Ck. RPM & Crit. Dimensions |  | 24 |  | DTM-C/V-#10
Conical Mold, Tamper | T84, C128 | Ck. Critical Dimensions |  | 24 |  | DTM-C/V-#12
Steel Balls | T96, C131 | Ck. Indiv. Wt. & Charge Wt. |  | 24 |  | DTM-C/V-#10
Sulfate Soundness Sample Containers | T104, C88 | Ck. Physical Condition |  | 12 |  | DTM-C/V-#14
Inspector's Signature
Title
Form DT-1695

## DT-1695A-Example
STATE OF TENNESSEE
DEPARTMENT OF TRANSPORTATION
DIVISION OF MATERIALS AND TESTS
6601 CENTENNIAL BLVD.
NASHVILLE, TENNESSEE 37243-0360
AGGREGATE TEST EQUIPMENT INSPECTION
[8 empty] | Date | 2003-01-12
Location | Clarksville Asphalt, Riverside Dr., Clarksville |  |  | County | MONTGOMERY |  |  | Region | 3
Inspector | Jim Williams
Equipment |  | Test Method | Requirement |  | Interval (Months) |  | Reference | Inspected | N/A
Bulk Density (Unit Weight) Measures |  | T19, C29 | Calibrate |  | 12 |  | AASHTO T19, ASTM C29
Mechanical Shakers |  |  | Ck. Sieving Thoroughness |  | 12 |  | AASHTO T27
Gen. Purpose Balances, Scales & Weights |  |  | Verify |  | 12 |  | AASHTO M231
Sieves, Coarse |  |  | Ck. Openings & Physical Condition |  | 6 |  | AASHTO M92, E11
Sieves, Fine |  |  | Ck. Physical Condition |  | 6 |  | AASHTO M92, E11
Ovens |  |  | Verify Temp. Settings |  | 4 |  | DMT-C/V-#3
Sulfate Oven |  | T104, C88 | Ck. Rate of Evaporation |  | 12 |  | AASHTO T104, DMT-C/V-#3
L.A. Machine |  | T96, C131 | Ck. RPM & Crit. Dimensions |  | 24 |  | DTM-C/V-#10
Conical Mold, Tamper |  | T84, C128 | Ck. Critical Dimensions |  | 24 |  | DTM-C/V-#12
Steel Balls |  | T96, C131 | Ck. Indiv. Wt. & Charge Wt. |  | 24 |  | DTM-C/V-#10
Sulfate Soundness Sample Containers |  | T104, C88 | Ck. Physical Condition |  | 12 |  | DTM-C/V-#14
Inspector's Signature
Transportation Technician II
Title
SAMPLE
Form DT-1695

## DT-1695B
STATE OF TENNESSEE
DEPARTMENT OF TRANSPORTATION
DIVISION OF MATERIALS AND TESTS
6601 CENTENNIAL BLVD.
NASHVILLE, TENNESSEE 37243-0360
BITUMINOUS MATERIALS TEST EQUIPMENT INSPECTION
Date
Location |  |  | County | 0 |  |  | Region | 0
Inspector
Equipment | Test Method | Requirement |  | Interval (Months) |  | Reference | Inspected | N/A
Saybolt Viscometers | T59, D244 | Calibrate |  | 36 |  | ASTM D88
Mechanical Shakers |  | Ck. Sieving Thoroughness |  | 12
Gen. Purpose Balances, Scales & Weights |  | Verify |  | 12 |  | AASHTO M231
Test Thermometers | T201, T202, T209, T228, T49, T51, T295, D5, D70, D113, D2041, D2170, D2171, D3142 | Calibrate |  | 6 |  | ASTM E1, E77
Analytical Balances & Weights |  | Verify |  | 24 |  | AASHTO M231
Comp. Test. Machine | T165, T167, T245, T246, D1074, D1075, D1559, D1560 | Verify Load Indications |  | 12 |  | ASTM E4
Mechanical Compactor | T245 | Calibrate |  | 36 |  | DMT-C/V-#25
CA Kneeding Compactor | T247, D1561 | Calibrate |  | 24 |  | AASHTO TP4
Timers | T201, T202, D2170, D2171 | Ck. Accuracy |  | 6 |  | DMT-C/V-#21
Ovens |  | Verify Temp. Settings |  | 4 |  | DMT-C/V-#3
Penetrometer and Accessories | T49, D5 | Ck. Dial & Timer Acc. & Needle Cond. |  | 6 |  | ASTM D5
Ductility Machine | T51, D113 | Ck. Molds & Travel Speed |  | 12
TFO & RTFO Oven | T179, T240, D1754, D2872 | Shelf/Carriage Ck. Rotation Speed & Temp. |  | 12 |  | ASTM E145
Vacuum System | T209, D2041 | Ck. Pressure |  | 12 |  | DMT-C/V-#29
Sieves, Coarse |  | Ck. Openings & Physical Condition |  | 6 |  | AASHTO M92, E11
Sieves, Fine |  | Ck. Physical Condition |  | 12 |  | AASHTO M92, E11
Molds, Followers, Calib. Cylinders | T246, T247, D1560, D1561 | Ck. Critical Dimensions |  | 12
Molds, Manual Comp. Hammers, Breaking Heads | T245, D1559 | Ck. Critical Dimensions & Mass of Hammer |  | 12 |  | DMT-C/V-#26, DMT-C/V-#27
Molds & Plungers | T167, D1074 | Ck. Critical Dimensions |  | 12
Brass Rings & Assembly | T53, D36 | Ck. Critical Dimensions |  | 12 |  | AASHTO T53
Pycnometers | T228, D70 | Ck. Critical Dimensions & Phys. Cond. |  | 12 |  | ASTM D70
Collars & Floats | T50, D139 | Ck. Critical Dimensions |  | 12 |  | ASTM D139
Flow Meters | T170, T240, D1856, D2872 | Calibrate |  | 12
Flash Cups | T48, T79, D92, D3143 | Ck. Critical Dimensions |  | 12 |  | ASTM D92
Inspector's Signature
Title
Form DT-1695

## DT-1695B-Example
STATE OF TENNESSEE
DEPARTMENT OF TRANSPORTATION
DIVISION OF MATERIALS AND TESTS
6601 CENTENNIAL BLVD.
NASHVILLE, TENNESSEE 37243-0360
BITUMINOUS MATERIALS TEST EQUIPMENT INSPECTION
[8 empty] | Date | 2003-01-12
Location | Clarksville Asphalt, Riverside Dr., Clarksville |  |  | County | MONTGOMERY |  |  | Region | 3
Inspector | Jim Williams
Equipment |  | Test Method | Requirement |  | Interval (Months) |  | Reference | Inspected | N/A
Saybolt Viscometers |  | T59, D244 | Calibrate |  | 36 |  | ASTM D88
Mechanical Shakers |  |  | Ck. Sieving Thoroughness |  | 12
Gen. Purpose Balances, Scales & Weights |  |  | Verify |  | 12 |  | AASHTO M231
Test Thermometers |  | T201, T202, T209, T228, T49, T51, T295, D5, D70, D113, D2041, D2170, D2171, D3142 | Calibrate |  | 6 |  | ASTM E1, E77
Analytical Balances & Weights |  |  | Verify |  | 24 |  | AASHTO M231
Comp. Test. Machine |  | T165, T167, T245, T246, D1074, D1075, D1559, D1560 | Verify Load Indications |  | 12 |  | ASTM E4
Mechanical Compactor |  | T245 | Calibrate |  | 36 |  | DMT-C/V-#25
CA Kneeding Compactor |  | T247, D1561 | Calibrate |  | 24 |  | AASHTO TP4
Timers |  | T201, T202, D2170, D2171 | Ck. Accuracy |  | 6 |  | DMT-C/V-#21
Ovens |  |  | Verify Temp. Settings |  | 4 |  | DMT-C/V-#3
Penetrometer and Accessories |  | T49, D5 | Ck. Dial & Timer Acc. & Needle Cond. |  | 6 |  | ASTM D5
Ductility Machine |  | T51, D113 | Ck. Molds & Travel Speed |  | 12
TFO & RTFO Oven |  | T179, T240, D1754, D2872 | Shelf/Carriage Ck. Rotation Speed & Temp. |  | 12 |  | ASTM E145
Vacuum System |  | T209, D2041 | Ck. Pressure |  | 12 |  | DMT-C/V-#29
Sieves, Coarse |  |  | Ck. Openings & Physical Condition |  | 6 |  | AASHTO M92, E11
Sieves, Fine |  |  | Ck. Physical Condition |  | 12 |  | AASHTO M92, E11
Molds, Followers, Calib. Cylinders |  | T246, T247, D1560, D1561 | Ck. Critical Dimensions |  | 12
Molds, Manual Comp. Hammers, Breaking Heads |  | T245, D1559 | Ck. Critical Dimensions & Mass of Hammer |  | 12 |  | DMT-C/V-#26, DMT-C/V-#27
Molds & Plungers |  | T167, D1074 | Ck. Critical Dimensions |  | 12
Brass Rings & Assembly |  | T53, D36 | Ck. Critical Dimensions |  | 12 |  | AASHTO T53
Pycnometers |  | T228, D70 | Ck. Critical Dimensions & Phys. Cond. |  | 12 |  | ASTM D70
Collars & Floats |  | T50, D139 | Ck. Critical Dimensions |  | 12 |  | ASTM D139
Flow Meters |  | T170, T240, D1856, D2872 | Calibrate |  | 12
Flash Cups |  | T48, T79, D92, D3143 | Ck. Critical Dimensions |  | 12 |  | ASTM D92
Inspector's Signature
Transportation Technician II
Title
SAMPLE
Form DT-1695

## DT-1695H
STATE OF TENNESSEE
DEPARTMENT OF TRANSPORTATION
DIVISION OF MATERIALS AND TESTS
6601 CENTENNIAL BLVD.
NASHVILLE, TENNESSEE 37243-0360
HYDRAULIC CEMENT TEST EQUIPMENT INSPECTION
Date
Location |  |  | County | 0 |  |  | Region | 0
Inspector
Equipment | Test Method | Requirement |  | Interval (Months) |  | Reference | Inspected | N/A
Moist Cabinet/Water Tank Recording Thermometer | T106, C109 | Check for Accuracy |  | 6 |  | AASHTO M201
Storage Water | T106, C109 | Check for Lime Saturation |  | 6
General Purpose Balances, Scales & Weights |  | Verify |  | 12 |  | AASHTO M231, ASTM C1005
Analytical Balances & Weights |  | Verify |  | 12 |  | AASHTO M231, ASTM C1005
No. 325 Sieves | T192, C430 | Clean after 5 Determinations, Calib. after 100 Determinations |  | ---- |  | AASHTO M92, ASTM E161
No. 325 Nozzle | T192, C430 | Check Flow Rate |  | 6 |  | AASHTO M92, ASTM E161
Comp. Testing Machine | T106, C109 | Verify Load Indications |  | 12
Bearing Blocks | T106, C109 | Check Planeness |  | 12
Wagner Turpidimeter | T98, C115 | Calibrate |  | 6 |  | AASHTO T98
Standard Sand |  | Each Shipment Check for Conformance to C778 |  | ---- |  | ASTM C778
Air-Permeability Apparatus | T153, C204 | Standardize |  | 30 |  | NIST 114
Flow Tables | M152, C230 | Verify Flow Results |  | 30 |  | AASHTO M152
Air Content Measures | T137, C185 | Calibrate |  | 30 |  | AASHTO T137
Cube Molds and Tampers | T106, C109 | Ck. Critical Dim. & Phys. Cond. |  | 30 |  | AASHTO T106
Vicat Apparatus and Vicat Ring | T129, T131, T186, C187, C191, C451 | Ck. Critical Dim., Phys. Cond., Mass |  | 30 |  | AASHTO T129
Gillmore Test Apparatus | T154, C266 | Ck. Critical Dim., Phys. Cond., Mass |  | 30 |  | AASHTO T154
Mechanical Mixing Apparatus | T162, C305 | Ck. Critical Clearances and Speeds |  | 30 |  | AASHTO T162
Water-Retention Apparatus | C91 | Ck. Critical Dimensions |  | 30 |  | ASTM C91
Inspector's Signature
Title
Form DT-1695

## DT-1695H-Example
STATE OF TENNESSEE
DEPARTMENT OF TRANSPORTATION
DIVISION OF MATERIALS AND TESTS
6601 CENTENNIAL BLVD.
NASHVILLE, TENNESSEE 37243-0360
HYDRAULIC CEMENT TEST EQUIPMENT INSPECTION
[8 empty] | Date | 2003-01-12
Location | Clarksville Asphalt, Riverside Dr., Clarksville |  |  | County | MONTGOMERY |  |  | Region | 3
Inspector | Jim Williams
Equipment |  | Test Method | Requirement |  | Interval (Months) |  | Reference | Inspected | N/A
Moist Cabinet/Water Tank Recording Thermometer |  | T106, C109 | Check for Accuracy |  | 6 |  | AASHTO M201
Storage Water |  | T106, C109 | Check for Lime Saturation |  | 6
General Purpose Balances, Scales & Weights |  |  | Verify |  | 12 |  | AASHTO M231, ASTM C1005
Analytical Balances & Weights |  |  | Verify |  | 12 |  | AASHTO M231, ASTM C1005
No. 325 Sieves |  | T192, C430 | Clean after 5 Determinations, Calib. after 100 Determinations |  | ---- |  | AASHTO M92, ASTM E161
No. 325 Nozzle |  | T192, C430 | Check Flow Rate |  | 6 |  | AASHTO M92, ASTM E161
Comp. Testing Machine |  | T106, C109 | Verify Load Indications |  | 12
Bearing Blocks |  | T106, C109 | Check Planeness |  | 12
Wagner Turpidimeter |  | T98, C115 | Calibrate |  | 6 |  | AASHTO T98
Standard Sand |  |  | Each Shipment Check for Conformance to C778 |  | ---- |  | ASTM C778
Air-Permeability Apparatus |  | T153, C204 | Standardize |  | 30 |  | NIST 114
Flow Tables |  | M152, C230 | Verify Flow Results |  | 30 |  | AASHTO M152
Air Content Measures |  | T137, C185 | Calibrate |  | 30 |  | AASHTO T137
Cube Molds and Tampers |  | T106, C109 | Ck. Critical Dim. & Phys. Cond. |  | 30 |  | AASHTO T106
Vicat Apparatus and Vicat Ring |  | T129, T131, T186, C187, C191, C451 | Ck. Critical Dim., Phys. Cond., Mass |  | 30 |  | AASHTO T129
Gillmore Test Apparatus |  | T154, C266 | Ck. Critical Dim., Phys. Cond., Mass |  | 30 |  | AASHTO T154
Mechanical Mixing Apparatus |  | T162, C305 | Ck. Critical Clearances and Speeds |  | 30 |  | AASHTO T162
Water-Retention Apparatus |  | C91 | Ck. Critical Dimensions |  | 30 |  | ASTM C91
Inspector's Signature
Transportation Technician II
Title
SAMPLE
Form DT-1695

## DT-1695P
STATE OF TENNESSEE
DEPARTMENT OF TRANSPORTATION
DIVISION OF MATERIALS AND TESTS
6601 CENTENNIAL BLVD.
NASHVILLE, TENNESSEE 37243-0360
PORTLAND CEMENT CONCRETE TEST EQUIPMENT INSPECTION
Date
Location |  |  | County | 0 |  |  | Region | 0
Inspector
Equipment | Test Method | Requirement |  | Interval (Months) |  | Reference | Inspected | N/A
Unit Weight Measures | T121, C138 | Calibrate |  | 12 |  | AASHTO T19, ASTM C29
Pressure Meter, Type B | T152, C173 | Calibrate |  | 3
Volumetric Air Meter | T196, C231 | Calibrate |  | 3
Comp. Test. Machine | T22, C39 | Verify Load Indications |  | 12 |  | ASTM E4
Capping Material |  | Check Strength |  | 3 |  | ASTM C1231
Slump Cones, Tamping Rods | T119, C143 | Ck. Critical Dimensions |  | 12 |  | DMT-C/V-#13
Single-Use Molds | T22, T23, C31, C39 | Ck. Dimensions of Each Shipment |  | ---- |  | AASHTO M205
Moist Room/Storage Tanks | C31, C39 | Temp. Verified with Recording Thermometer |  | ---- |  | AASHTO M201
Gen. Purpose Balances, Scales & Weights |  | Verify |  | 12 |  | AASHTO M231
Inspector's Signature
Title
Form DT-1695

## DT-1695P-Example
STATE OF TENNESSEE
DEPARTMENT OF TRANSPORTATION
DIVISION OF MATERIALS AND TESTS
6601 CENTENNIAL BLVD.
NASHVILLE, TENNESSEE 37243-0360
PORTLAND CEMENT CONCRETE TEST EQUIPMENT INSPECTION
[8 empty] | Date | 2003-01-12
Location | Clarksville Asphalt, Riverside Dr., Clarksville |  |  | County | MONTGOMERY |  |  | Region | 3
Inspector | Jim Williams
Equipment |  | Test Method | Requirement |  | Interval (Months) |  | Reference | Inspected | N/A
Unit Weight Measures |  | T121, C138 | Calibrate |  | 12 |  | AASHTO T19, ASTM C29
Pressure Meter, Type B |  | T152, C173 | Calibrate |  | 3
Volumetric Air Meter |  | T196, C231 | Calibrate |  | 3
Comp. Test. Machine |  | T22, C39 | Verify Load Indications |  | 12 |  | ASTM E4
Capping Material |  |  | Check Strength |  | 3 |  | ASTM C1231
Slump Cones, Tamping Rods |  | T119, C143 | Ck. Critical Dimensions |  | 12 |  | DMT-C/V-#13
Single-Use Molds |  | T22, T23, C31, C39 | Ck. Dimensions of Each Shipment |  | ---- |  | AASHTO M205
Moist Room/Storage Tanks |  | C31, C39 | Temp. Verified with Recording Thermometer |  | ---- |  | AASHTO M201
Gen. Purpose Balances, Scales & Weights |  |  | Verify |  | 12 |  | AASHTO M231
Inspector's Signature
Transportation Technician II
Title
SAMPLE
Form DT-1695

## DT-1695S
STATE OF TENNESSEE
DEPARTMENT OF TRANSPORTATION
DIVISION OF MATERIALS AND TESTS
6601 CENTENNIAL BLVD.
NASHVILLE, TENNESSEE 37243-0360
SOILS TEST EQUIPMENT INSPECTION
Date
Location |  |  | County | 0 |  |  | Region | 0
Inspector
Equipment | Test Method | Requirement |  | Interval (Months) |  | Reference | Inspected | N/A
Mechanical Shakers |  | Ck. Sieving Thoroughness |  | 12
Gen. Purpose Balances, Scales & Weights |  | Verify |  | 12 |  | AASHTO M231
Compression or Loading Device | T193, T208, T216, T236, T296, T297, D1883, D2166, D2435, D2850, D3080, D4767 | Verify Load Indications |  | 12 |  | ASTM E4
Mechanical Compactor | T99, T180, D698, D1557 | Calibrate |  | 12 |  | DMT-C/V-#25
CA Kneeding Compactor | T190, D2844 | Calibrate |  | 24 |  | AASHTO TP4
Ovens |  | Verify Temp. Settings |  | 4 |  | DMT-C/V-#3
Vacuum System | T100, D854 | Ck. Pressure |  | 24 |  | DMT-C/V-#29
Molds | T99, T134, T135, T136, T180, T190, T193, D698, D558, D559, D560, D1557, D1883, D2844 | Ck. Critical Dimensions |  | 12
Manual Hammer | T99, T180, D698, D1557 | Ck. Weight & Critical Dimensions |  | 12
Sieves, Coarse |  | Ck. Openings & Physical Condition |  | 6 |  | AASHTO M92, E11
Sieves, Fine |  | Ck. Physical Condition |  | 6 |  | AASHTO M92, E11
Liquid Limit Device | T89, D4318 | Ck. Wear & Critical Dimensions |  | 12
Grooving Tool | T89, D4318 | Ck. Critical Dimensions |  | 12
Hydrometers | T88, D422 | Ck. Critical Dimensions |  | 24
Straightedge | T99, T134, T135, T136, T180, D698, D558, D559, D560, D1557 | Ck. Planeness of Edge |  | 6
Weighted Foot Assembly | T176, D2419 | Ck. Weight |  | 12
CBR Annular and Slotted Weights | T193, D1883 | Ck. Diameter |  | 12
CBR Penetration Piston | T193, D1883 | Ck. Outside Diameter |  | 12
Standard Metal Specimen | T190, D2844 | Ck. Diameter |  | 12
Metal Follower | T190, D2844 |  |  | 12
Inspector's Signature
Title
Form DT-1695

## DT-1695S-Example
STATE OF TENNESSEE
DEPARTMENT OF TRANSPORTATION
DIVISION OF MATERIALS AND TESTS
6601 CENTENNIAL BLVD.
NASHVILLE, TENNESSEE 37243-0360
SOILS TEST EQUIPMENT INSPECTION
[8 empty] | Date | 2003-01-12
Location | Clarksville Asphalt, Riverside Dr., Clarksville |  |  | County | MONTGOMERY |  |  | Region | 3
Inspector | Jim Williams
Equipment |  | Test Method | Requirement |  | Interval (Months) |  | Reference | Inspected | N/A
Mechanical Shakers |  |  | Ck. Sieving Thoroughness |  | 12
Gen. Purpose Balances, Scales & Weights |  |  | Verify |  | 12 |  | AASHTO M231
Compression or Loading Device |  | T193, T208, T216, T236, T296, T297, D1883, D2166, D2435, D2850, D3080, D4767 | Verify Load Indications |  | 12 |  | ASTM E4
Mechanical Compactor |  | T99, T180, D698, D1557 | Calibrate |  | 12 |  | DMT-C/V-#25
CA Kneeding Compactor |  | T190, D2844 | Calibrate |  | 24 |  | AASHTO TP4
Ovens |  |  | Verify Temp. Settings |  | 4 |  | DMT-C/V-#3
Vacuum System |  | T100, D854 | Ck. Pressure |  | 24 |  | DMT-C/V-#29
Molds |  | T99, T134, T135, T136, T180, T190, T193, D698, D558, D559, D560, D1557, D1883, D2844 | Ck. Critical Dimensions |  | 12
Manual Hammer |  | T99, T180, D698, D1557 | Ck. Weight & Critical Dimensions |  | 12
Sieves, Coarse |  |  | Ck. Openings & Physical Condition |  | 6 |  | AASHTO M92, E11
Sieves, Fine |  |  | Ck. Physical Condition |  | 6 |  | AASHTO M92, E11
Liquid Limit Device |  | T89, D4318 | Ck. Wear & Critical Dimensions |  | 12
Grooving Tool |  | T89, D4318 | Ck. Critical Dimensions |  | 12
Hydrometers |  | T88, D422 | Ck. Critical Dimensions |  | 24
Straightedge |  | T99, T134, T135, T136, T180, D698, D558, D559, D560, D1557 | Ck. Planeness of Edge |  | 6
Weighted Foot Assembly |  | T176, D2419 | Ck. Weight |  | 12
CBR Annular and Slotted Weights |  | T193, D1883 | Ck. Diameter |  | 12
CBR Penetration Piston |  | T193, D1883 | Ck. Outside Diameter |  | 12
Standard Metal Specimen |  | T190, D2844 | Ck. Diameter |  | 12
Metal Follower |  | T190, D2844 |  |  | 12
Inspector's Signature
Transportation Technician II
Title
SAMPLE
Form DT-1695

## Balances
LABORATORY INSPECTION SHEET
BALANCES
Date
Location |  |  | County | 0 |  |  | Region | 0
Inspector
ID or Serial No. |  |  |  |  | Capacity
Weight | Reading | % Error | Notes
ID or Serial No. |  |  |  |  | Capacity
Weight | Reading | % Error | Notes
ID or Serial No. |  |  |  |  | Capacity
Weight | Reading | % Error | Notes
ID or Serial No. |  |  |  |  | Capacity
Weight | Reading | % Error | Notes
Remarks:
Inspector's Signature
Title

## Balances Example
LABORATORY INSPECTION SHEET
BALANCES
[10 empty] | Date | 2003-01-12
Location | Clarksville Asphalt, Riverside Dr., Clarksville |  |  |  | County | MONTGOMERY |  |  |  | Region | 3
Inspector | Jim Williams
ID or Serial No. |  | 54136 |  |  |  |  | Capacity |  | 5000 g
Weight |  |  | Reading | % Error | Notes
500 |  |  | 500 | 0
1000 |  |  | 1000 | 0
1500 |  |  | 1500 | 0
2000 |  |  | 2000 | 0
3000 |  |  | 3000 | 0
4000 |  |  | 3999.5 | -0.0125
5000 |  |  | 4999 | -0.02
ID or Serial No. |  | 7867 |  |  |  |  | Capacity |  | 5000 g
Weight |  |  | Reading | % Error | Notes
500 |  |  | 500 | 0
1000 |  |  | 1000 | 0
1500 |  |  | 1500 | 0
2000 |  |  | 2000 | 0
3000 |  |  | 3000 | 0
4000 |  |  | 3999.7 | -0.0075
5000 |  |  | 4999 | -0.02
ID or Serial No. |  | 36876 |  |  |  |  | Capacity |  | 5000 g
Weight |  |  | Reading | % Error | Notes
500 |  |  | 500 | 0
1000 |  |  | 1000 | 0
1500 |  |  | 1500 | 0
2000 |  |  | 2000 | 0
3000 |  |  | 3000 | 0
4000 |  |  | 3998 | -0.05
5000 |  |  | 4996 | -0.08
ID or Serial No. |  | 38766 |  |  |  |  | Capacity |  | 5000 g
Weight |  |  | Reading | % Error | Notes
500 |  |  | 500 | 0
1000 |  |  | 1000 | 0
1500 |  |  | 1500 | 0
2000 |  |  | 2000 | 0
3000 |  |  | 3000 | 0
4000 |  |  | 3997 | -0.075
5000 |  |  | 4996 | -0.08
Remarks:
Inspector's Signature
Transportation Technician II
Title
SAMPLE

## Marshall Molds
LABORATORY INSPECTION SHEET
MARSHALL MOLDS
Date
Location |  |  |  | County | 0 |  |  | Region | 0
Inspector
Mold Number | Inside Diameter |  |  |  |  | Notes
| 1 | 2 | 3 | 4
Remarks:
Inspector's Signature
Title

## Marshall Molds Example
LABORATORY INSPECTION SHEET
MARSHALL MOLDS
[9 empty] | Date | 2003-01-12
Location | Clarksville Asphalt, Riverside Dr., Clarksville |  |  |  | County | MONTGOMERY |  |  | Region | 3
Inspector | Jim Williams
Mold Number |  | Inside Diameter |  |  |  |  | Notes
|  | 1 | 2 | 3 | 4
12 |  | 4.5 | 4.5 | 4.495 | 4.49
13 |  | 4.5 | 4.49 | 4.5 | 4.5
14 |  | 4.5 | 4.5 | 4.49 | 4.5
16 |  | 4.5 | 4.5 | 4.5 | 4.49
19 |  | 4.49 | 4.5 | 4.5 | 4.5
Remarks:
Inspector's Signature
Transportation Technician II
Title
SAMPLE

## Ovens
LABORATORY INSPECTION SHEET
OVENS
Date
Location |  |  | County |  | 0 |  |  | Region | 0
Inspector
ID or Serial No. | [5 empty] | Specified Temp.
Time |  | Temperature |  | Notes
Begin | End
ID or Serial No. | [5 empty] | Specified Temp.
Time |  | Temperature |  | Notes
Begin | End
ID or Serial No. | [5 empty] | Specified Temp.
Time |  | Temperature |  | Notes
Begin | End
ID or Serial No. | [5 empty] | Specified Temp.
Time |  | Temperature |  | Notes
Begin | End
Remarks:
Inspector's Signature
Title

## Ovens Example
LABORATORY INSPECTION SHEET
OVENS
[11 empty] | Date | 2003-01-12
Location | Clarksville Asphalt, Riverside Dr., Clarksville |  |  |  | County |  | MONTGOMERY |  |  |  | Region | 3
Inspector | Jim Williams
ID or Serial No. |  | 453125A | [5 empty] | Specified Temp. |  | 230
Time |  |  |  | Temperature |  | Notes
Begin |  |  | End
12:00 |  |  |  | 230 |  | oven closed; set at specified temperature
13:00 |  |  |  | 231 |  | first reading
13:30 |  |  |  | 230 |  | second reading
14:00 |  |  |  | 229 |  | third reading
|  |  | 14:30 | 230 |  | final reading
ID or Serial No. |  | 453125B | [5 empty] | Specified Temp. |  | 230
Time |  |  |  | Temperature |  | Notes
Begin |  |  | End
12:30 |  |  |  | 230 |  | oven closed; set at specified temperature
13:30 |  |  |  | 231 |  | first reading
14:00 |  |  |  | 232 |  | second reading
14:30 |  |  |  | 233 |  | third reading
|  |  | 15:00 | 230 |  | final reading
ID or Serial No. |  | 453125C | [5 empty] | Specified Temp. |  | 230
Time |  |  |  | Temperature |  | Notes
Begin |  |  | End
13:00 |  |  |  | 230 |  | oven closed; set at specified temperature
14:00 |  |  |  | 231 |  | first reading
14:30 |  |  |  | 232 |  | second reading
15:00 |  |  |  | 233 |  | third reading
|  |  | 15:30 | 230 |  | final reading
ID or Serial No. |  | 453125D | [5 empty] | Specified Temp. |  | 230
Time |  |  |  | Temperature |  | Notes
Begin |  |  | End
13:30 |  |  |  | 230 |  | oven closed; set at specified temperature
14:30 |  |  |  | 233 |  | first reading
15:00 |  |  |  | 235 |  | second reading
15:30 |  |  |  | 229 |  | third reading
|  |  | 16:00 | 230 |  | final reading
Remarks:
Inspector's Signature
Transportation Technician II
Title
All of the ovens above are used solely for drying of aggregates.
SAMPLE

## Thermometers
LABORATORY INSPECTION SHEET
THERMOMETERS
Date
Location |  |  | County |  | 0 |  |  | Region | 0
Inspector
ID or Serial No. | Temperature Setting | Actual Reading |  | Correction |  |  | Notes
Remarks:
Inspector's Signature
Title

## Thermometers Example
LABORATORY INSPECTION SHEET
THERMOMETERS
[9 empty] | Date | 2003-01-12
Location | Clarksville Asphalt, Riverside Dr., Clarksville |  |  | County |  | MONTGOMERY |  |  | Region | 3
Inspector | Jim Williams
ID or Serial No. |  | Temperature Setting | Actual Reading |  | Correction |  |  | Notes
3425A |  | 150 | 149.6 |  | 0.4
|  | 170 | 169.44 |  | 0.56
|  | 190 | 189.41 |  | 0.59
|  | 210 | 209.3 |  | 0.7
|  | 230 | 229.22 |  | 0.78
|  | 250 | 249.13 |  | 0.87
3426A |  | 150 | 150 |  | 0
|  | 170 | 169.55 |  | 0.45
|  | 190 | 189.46 |  | 0.54
|  | 210 | 209.44 |  | 0.56
|  | 230 | 229.25 |  | 0.75
|  | 250 | 249.15 |  | 0.85
Remarks:
Inspector's Signature
Transportation Technician II
Title
SAMPLE

## Water Baths
LABORATORY INSPECTION SHEET
WATER BATHS
Date
Location |  |  | County |  | 0 |  |  | Region | 0
Inspector
ID or Serial No. | Temperature Setting | Actual Reading |  | Correction |  |  | Notes
Remarks:
Inspector's Signature
Title

## Water Baths Example
LABORATORY INSPECTION SHEET
WATER BATHS
[9 empty] | Date | 2003-01-12
Location | Clarksville Asphalt, Riverside Dr., Clarksville |  |  | County |  | MONTGOMERY |  |  | Region | 3
Inspector | Jim Williams
ID or Serial No. |  | Temperature Setting | Actual Reading |  | Correction |  |  | Notes
23 |  | 100 | 100 |  | 0
"" |  | 120 | 119.68 |  | 0.32
"" |  | 140 | 139.5 |  | 0.5
"" |  | 160 | 159.45 |  | 0.55
"" |  | 180 | 179.34 |  | 0.66
"" |  | 200 | 199.22 |  | 0.78
"" |  | 220 | 219.15 |  | 0.85
"" |  | 240 | 239.06 |  | 0.94
"" |  | 260 | 259 |  | 1
"" |  | 280 | 278.98 |  | 1.02
"" |  | 300 | 298.66 |  | 1.34
"" |  | 320 | 318.53 |  | 1.47
"" |  | 340 | 338.42 |  | 1.58
"" |  | 360 | 358.21 |  | 1.79
"" |  | 380 | 378.1 |  | 1.9
"" |  | 400 | 398 |  | 2
Remarks:
Inspector's Signature
Transportation Technician II
Title
SAMPLE

##  (hidden)
cell range for combo box |  |  |  |  | cell link
| County | Region | County Number |  | 96 | 96 | 96 | 96 | 96
1 | ANDERSON | 1 | 01 | ANDERSON | 90
2 | BEDFORD | 3 | 02 | BEDFORD
3 | BENTON | 4 | 03 | BENTON
4 | BLEDSOE | 2 | 04 | BLEDSOE
5 | BLOUNT | 1 | 05 | BLOUNT
6 | BRADLEY | 2 | 06 | BRADLEY
7 | CAMPBELL | 1 | 07 | CAMPBELL
8 | CANNON | 2 | 08 | CANNON
9 | CARROLL | 4 | 09 | CARROLL
10 | CARTER | 1 | 10 | CARTER
11 | CHEATHAM | 3 | 11 | CHEATHAM
12 | CHESTER | 4 | 12 | CHESTER
13 | CLAIBORNE | 1 | 13 | CLAIBORNE
14 | CLAY | 2 | 14 | CLAY
15 | COCKE | 1 | 15 | COCKE
16 | COFFEE | 2 | 16 | COFFEE
17 | CROCKETT | 4 | 17 | CROCKETT
18 | CUMBERLAND | 2 | 18 | CUMBERLAND
19 | DAVIDSON | 3 | 19 | DAVIDSON
20 | DECATUR | 4 | 20 | DECATUR
21 | DEKALB | 2 | 21 | DEKALB
22 | DICKSON | 3 | 22 | DICKSON
23 | DYER | 4 | 23 | DYER
24 | FAYETTE | 4 | 24 | FAYETTE
25 | FENTRESS | 2 | 25 | FENTRESS
26 | FRANKLIN | 2 | 26 | FRANKLIN
27 | GIBSON | 4 | 27 | GIBSON
28 | GILES | 3 | 28 | GILES
29 | GRAINGER | 1 | 29 | GRAINGER
30 | GREENE | 1 | 30 | GREENE
31 | GRUNDY | 2 | 31 | GRUNDY
32 | HAMBLEN | 1 | 32 | HAMBLEN
33 | HAMILTON | 2 | 33 | HAMILTON
34 | HANCOCK | 1 | 34 | HANCOCK
35 | HARDEMAN | 4 | 35 | HARDEMAN
36 | HARDIN | 4 | 36 | HARDIN
37 | HAWKINS | 1 | 37 | HAWKINS
38 | HAYWOOD | 4 | 38 | HAYWOOD
39 | HENDERSON | 4 | 39 | HENDERSON
40 | HENRY | 4 | 40 | HENRY
41 | HICKMAN | 3 | 41 | HICKMAN
42 | HOUSTON | 3 | 42 | HOUSTON
43 | HUMPHREY | 3 | 43 | HUMPHREY
44 | JACKSON | 2 | 44 | JACKSON
45 | JEFFERSON | 1 | 45 | JEFFERSON
46 | JOHNSON | 1 | 46 | JOHNSON
47 | KNOX | 1 | 47 | KNOX
48 | LAKE | 4 | 48 | LAKE
49 | LAUDERDALE | 4 | 49 | LAUDERDALE
50 | LAWRENCE | 3 | 50 | LAWRENCE
51 | LEWIS | 3 | 51 | LEWIS
52 | LINCOLN | 3 | 52 | LINCOLN
53 | LOUDON | 1 | 53 | LOUDON
54 | MACON | 3 | 56 | MACON
55 | MADISON | 4 | 57 | MADISON
56 | MARION | 2 | 58 | MARION
57 | MARSHALL | 3 | 59 | MARSHALL
58 | MAURY | 3 | 60 | MAURY
59 | MCMINN | 2 | 54 | MCMINN
60 | MCNAIRY | 4 | 55 | MCNAIRY
61 | MEIGS | 2 | 61 | MEIGS
62 | MONROE | 1 | 62 | MONROE
63 | MONTGOMERY | 3 | 63 | MONTGOMERY
64 | MOORE | 3 | 64 | MOORE
65 | MORGAN | 1 | 65 | MORGAN
66 | OBION | 4 | 66 | OBION
67 | OVERTON | 2 | 67 | OVERTON
68 | PERRY | 3 | 68 | PERRY
69 | PICKETT | 2 | 69 | PICKETT
70 | POLK | 2 | 70 | POLK
71 | PUTNAM | 2 | 71 | PUTNAM
72 | RHEA | 2 | 72 | RHEA
73 | ROANE | 1 | 73 | ROANE
74 | ROBERTSON | 3 | 74 | ROBERTSON
75 | RUTHERFORD | 3 | 75 | RUTHERFORD
76 | SCOTT | 1 | 76 | SCOTT
77 | SEQUATCHIE | 2 | 77 | SEQUATCHIE
78 | SEVIER | 1 | 78 | SEVIER
79 | SHELBY | 4 | 79 | SHELBY
80 | SMITH | 3 | 80 | SMITH
81 | STEWART | 3 | 81 | STEWART
82 | SULLIVAN | 1 | 82 | SULLIVAN
83 | SUMNER | 3 | 83 | SUMNER
84 | TIPTON | 4 | 84 | TIPTON
85 | TROUSDALE | 3 | 85 | TROUSDALE
86 | UNICOI | 1 | 86 | UNICOI
87 | UNION | 1 | 87 | UNION
88 | VAN BUREN | 2 | 88 | VAN BUREN
89 | WARREN | 2 | 89 | WARREN
90 | WASHINGTON | 1 | 90 | WASHINGTON
91 | WAYNE | 3 | 91 | WAYNE
92 | WEAKLEY | 4 | 92 | WEAKLEY
93 | WHITE | 2 | 93 | WHITE
94 | WILLIAMSON | 3 | 94 | WILLIAMSON
95 | WILSON | 3 | 95 | WILSON
96 | (Leave blank)